## Sheet1
iBeacon1 Packet
Advertisement parameters: Tx Power - Configured by manufacturer to meet 70 m broadcast distance ADV Interval - 10 sec
Byte Number | Value | Description | Remark | Parameter description (GATT connection)
00 | 0x02 | Data length - 29bytes
01 | 0x01 | Data type
02 | 0x06 | Value
03 | 0x1A | Length of the following section
04 | 0xFF | Custom Manufacturer Data
05 - 06 | 0x004C | Apple's Bluetooth SIG Code
07 | 0x02 | Apple's iBeacon Advertising Indicator
08 | 0x15 | UUID + Major + Minor + TXPower
09 - 24 | 0x00000000000000000000000000000000 to 0xFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFF | Proximity UUID | Used as the UUID for the device. | Property: Read/Write
25 - 26 | 0x0000 to 0xFFFF | Major | Identifier to distinguish this iBeacon device | Property: Read/Write
27 | 0x00 to 0xFF | Minor | Identifier to distinguish this iBeacon device
28 | 0b0 to 0b1 | iBeacon Indicator | 0 for the first iBeacon packet 1 for the second
28 | 0b0 to 0b1 | Battery Indicator | Indicates a 0 if the unit is under 25% battery life(Red State) and a 1 above 25% (Yellow or Green.)
28 | 0b000000 to 0b111111 | Time since Last Sterilization | Uses the sequence on the Last Sterilize Decoder to show the last sterilization up to 30 days. Only reset when the tag goes over temperature. |  | Property: Read/Write
29 | 0x00 to 0xFF | TXPower | Strength of the signal 1 meter away from the Beacon in 2's complement | Property: Read only
iBeacon2 Packet
Advertisement parameters: Tx Power - Configured by manufacturer to meet 70 m broadcast distance ADV Interval - 10 sec
Byte Number | Value | Description | Remark | Parameter description (GATT connection)
00 | 0x02 | Data length - 29bytes
01 | 0x01 | Data type
02 | 0x06 | Value
03 | 0x1A | Length of the following section
04 | 0xFF | Custom Manufacturer Data
05 - 06 | 0x004C | Apple's Bluetooth SIG Code
07 | 0x02 | Apple's iBeacon Advertising Indicator
08 | 0x15 | UUID + Major + Minor + TXPower
09 - 24 | 0x00000000000000000000000000000000 to 0xFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFFF | Proximity UUID | Used as the UUID for the device. | Property: Read/Write
25 - 26 | 0x0000 to 0xFFFF | Major | Identifier to distinguish this iBeacon device | Property: Read/Write
27 | 0x00 to 0xFF | Minor | Identifier to distinguish this iBeacon device
28 | 0b0 to 0b1 | iBeacon Indicator | 0 for the first iBeacon packet 1 for the second | Property: Read/Write
28 | 0b0 to 0b1 | Second Battery Indicator | This will indicate whether the battery is between 100 to 75% (Green State) with 1 and below 75% (Yellow State)
28 | 0b000000 to 0b111111 | Total Number of Autoclave Cycles | Count the number of autoclave cycles to 63 and then roll over. |  | Property: Read/Write
29 | 0x00 to 0xFF | TXPower | Strength of the signal 1 meter away from the Beacon in 2's complement | Property: Read only

## Lifetime
Tag Value | Number of Cycles
0 | 0
1 | 1
2 | 2
3 | 3
4 | 4
5 | 5
6 | 6
7 | 7
8 | 8
9 | 9
10 | 10
11 | 11
12 | 12
13 | 13
14 | 14
15 | 15
16 | 16
17 | 17
18 | 18
19 | 19
20 | 20
21 | 21
22 | 22
23 | 23
24 | 24
25 | 25
26 | 26
27 | 27
28 | 28
29 | 29
30 | 30
31 | 31
32 | 32
33 | 33
34 | 34
35 | 35
36 | 36
37 | 37
38 | 38
39 | 39
40 | 40
41 | 41
42 | 42
43 | 43
44 | 44
45 | 45
46 | 46
47 | 47
48 | 48
49 | 49
50 | 50
51 | 51
52 | 52
53 | 53
54 | 54
55 | 55
56 | 56
57 | 57
58 | 58
59 | 59
60 | 60
61 | 61
62 | 62
63 | 63

## Last Sterilize Decoder
STT Tag Value | 0 | Increment（单位） | Min Hours | Max Hours | Min Days | Max Days | months
0 | 1 |  | 0 | 1 | 0 | 0.041667 | 0.0013889
1 | 2 | 1 | 1 | 2 | 0.041667 | 0.083333 | 0.0027778
2 | 3 | 1 | 2 | 3 | 0.083333 | 0.125 | 0.0041667
3 | 4 | 1 | 3 | 4 | 0.125 | 0.16667 | 0.0055556
4 | 5 | 1 | 4 | 5 | 0.16667 | 0.20833 | 0.0069444
5 | 6 | 1 | 5 | 6 | 0.20833 | 0.25 | 0.0083333
6 | 7 | 1 | 6 | 7 | 0.25 | 0.29167 | 0.0097222
7 | 8 | 1 | 7 | 8 | 0.29167 | 0.33333 | 0.011111
8 | 9 | 1 | 8 | 9 | 0.33333 | 0.375 | 0.0125
9 | 10 | 1 | 9 | 10 | 0.375 | 0.41667 | 0.013889
10 | 11 | 1 | 10 | 11 | 0.41667 | 0.45833 | 0.015278
11 | 12 | 1 | 11 | 12 | 0.45833 | 0.5 | 0.016667
12 | 24 | 12 |  | 13 | 0.5 | 1 | 0.033333 | 1单位 | 每增加1 就加1
13 | 48 | 24 | 24 | 14 | 1 | 2 | 0.066667
14 | 72 | 24 | 48 | 15 | 2 | 3 | 0.1
15 | 96 | 24 | 72 | 16 | 3 | 4 | 0.13333
16 | 120 | 24 | 96 | 17 | 4 | 5 | 0.16667
17 | 144 | 24 | 120 | 18 | 5 | 6 | 0.2
18 | 168 | 24 | 144 | 19 | 6 | 7 | 0.23333
19 | 192 | 24 | 168 | 20 | 7 | 8 | 0.26667
20 | 216 | 24 | 192 | 21 | 8 | 9 | 0.3
21 | 240 | 24 | 216 | 22 | 9 | 10 | 0.33333
22 | 264 | 24 | 240 | 23 | 10 | 11 | 0.36667
23 | 288 | 24 | 264 | 24 | 11 | 12 | 0.4 | 每
24 | 312 | 24 | 288 | 25 | 12 | 13 | 0.43333 | 加
25 | 336 | 24 | 312 | 26 | 13 | 14 | 0.46667 | 24
26 | 360 | 24 | 336 | 27 | 14 | 15 | 0.5 | 就
27 | 384 | 24 | 360 | 28 | 15 | 16 | 0.53333 | 加
28 | 408 | 24 | 384 | 29 | 16 | 17 | 0.56667 | 1
29 | 432 | 24 | 408 | 30 | 17 | 18 | 0.6
30 | 456 | 24 | 432 | 31 | 18 | 19 | 0.63333
31 | 480 | 24 | 456 | 32 | 19 | 20 | 0.66667
32 | 504 | 24 | 480 | 33 | 20 | 21 | 0.7
33 | 528 | 24 | 504 | 34 | 21 | 22 | 0.73333
34 | 552 | 24 | 528 | 35 | 22 | 23 | 0.76667
35 | 576 | 24 | 552 | 36 | 23 | 24 | 0.8
36 | 600 | 24 | 576 | 37 | 24 | 25 | 0.83333
37 | 624 | 24 | 600 | 38 | 25 | 26 | 0.86667
38 | 648 | 24 | 624 | 39 | 26 | 27 | 0.9
39 | 672 | 24 | 648 | 40 | 27 | 28 | 0.93333
40 | 696 | 24 | 672 | 41 | 28 | 29 | 0.96667
41 | 720 | 24 | 696 | 42 | 29 | 30 | 1
42 | 744 | 24 | 720 | 43 | 30 | 31 | 1.0333 | 24单位
43 | 912 | 168 | 744 | 44 | 31 | 38 | 1.2667
44 | 1080 | 168 | 912 | 45 | 38 | 45 | 1.5
45 | 1248 | 168 | 1080 | 46 | 45 | 52 | 1.7333
46 | 1416 | 168 | 1248 | 47 | 52 | 59 | 1.9667
47 | 1584 | 168 | 1416 | 48 | 59 | 66 | 2.2
48 | 1752 | 168 | 1584 | 49 | 66 | 73 | 2.4333
49 | 1920 | 168 | 1752 | 50 | 73 | 80 | 2.6667
50 | 2088 | 168 | 1920 | 51 | 80 | 87 | 2.9
51 | 2256 | 168 | 2088 | 52 | 87 | 94 | 3.1333
52 | 2424 | 168 | 2256 | 53 | 94 | 101 | 3.3667
53 | 2592 | 168 | 2424 | 54 | 101 | 108 | 3.6
54 | 2760 | 168 | 2592 | 55 | 108 | 115 | 3.8333
55 | 2928 | 168 | 2760 | 56 | 115 | 122 | 4.0667
56 | 3096 | 168 | 2928 | 57 | 122 | 129 | 4.3
57 | 3264 | 168 | 3096 | 58 | 129 | 136 | 4.5333
58 | 3432 | 168 | 3264 | 59 | 136 | 143 | 4.7667
59 | 3600 | 168 | 3432 | 60 | 143 | 150 | 5
60 | 3768 | 168 | 3600 | 61 | 150 | 157 | 5.2333
61 | 3936 | 168 | 3768 | 62 | 157 | 164 | 5.4667
62 | 4104 | 168 | 3936 | 63 | 164 | 171 | 5.7
63 | 4272 | 168 | 4104 | 64 | 171 | 178 | 5.9333 | 111111b

## Battery Decoder
IBecaon Packet Indicator | Battery Indicator | Battery State
1 | 1 | Green (>75%)
1 | 0 | Yellow (>25% and <75%)
0 | 1 | Yellow or Green
0 | 0 | Red (<25%)