## Design Requirements
Set Track and Trace Requirements Traceability Matrix
# | Major Objective | User Needs | Design Input | Translations | Design Output | Notes | Test ID
1.01 | General System Requirements | Unit should fit within surgical trays and minimize size. | Unit profile should be equal to or less than 57x32x12mm. | 单位外形应等于或小于57x32x11mm。
1.02 |  | Unit should minimize weight. | Weight should be less than 20 grams. | 重量应少于20克。
1.03 |  | Unit should maximize lifespan in the field. | Unit can withstand 750 power cycles (sterilization cycles). | 本机可承受750次功率循环（灭菌循环）
1.04 |  | Unit should have extended battery life to minimize field maintenance requirements. | Installed battery should provide 4 years lifespan on least frequent broadcast settings. | 广播不频繁情况下 4年使用寿命
1.05 |  | Unit should comply with regulatory certifications. | Design until to comply with CE (EN 62368, EN 300328, EN 62311, EN301489-1, EN 301489-17, EN 55032, EN55035), FCC (FCC part15B, FCC part15C, FCC-ID). | 设计需要符合 CE（EN 62368、EN 300328、EN 62311、EN301489-1、EN 301489-17、EN 55032、EN55035）、FCC（FCC part15B、FCC part15C、FCC-ID）。
2.01 | Mechancial Requirements | The unit case should be durable and dark in color to prevent/mask aesthetic wear and tear. | Casing should be dark gray to black, made of resilient plastic such as PPSU with gaskets/o-rings for sealing made of appropriate material such as EPDM. | 外壳应为深灰色至黑色，由弹性塑料(如PPSU)制成，密封垫片/o形圈由适当材料(如三元乙丙橡胶EPDM)制成。
2.02 |  | The unit should be fixable to the surgical tray in at least two locations by zip ties. | Two holes for zip tie application should be included on opposite sides of the casing design. | 两个孔用于zip tie应用应包括在套管设计的反面。
2.03 |  | Battery should be capable of withstanding autoclave sterilization. |  | 一次性电池，化学成分，如二氧化锰(有待确认)
2.04 |  | Unit should be able to withstand repeated drops experienced over years of handling at a hospital. | Comply with standard JESD22-B111A or similar standard. | 符合JESD22-B111A（手持电子产品组件的板级跌落试验）标准或类似标准。
2.05 |  | Unit shall be able to operate between -25 C and 125 C | Unit and electronics shall be designed to be functional within this temperature range | 装置和电子器件的设计应在此温度范围内正常工作 -25 C ~ 125 C
2.06 |  | Unit should be corrosion resistant. | No physical changes should be observed when exposed to the following: 168h Motor oil 48h Salt water (salinity 10%) 48h Sulficirc acid (10%, pH 2) 48h NaOH (10%, pH 13) | 当接触以下物品时，不应观察到任何物理变化: 168小时 机油 48小时 海水(含盐量10%) 48小时 硫酸 (10%， pH 2) 48小时 氢氧化钠 (10%， pH 13)
2.07 |  | Unit should be resistant dust from standard environemntal conditions and to submersion in water that could accur during cleaning. | Unit should meet IP68 requirements or higher. | 机组应符合IP68或更高的要求。
2.08 |  | Unit should be resistant to normal vibration caused by shipping/handling and hospital usage. | Comply with standard JESD22-B103B.01 or similar standard. | 符合JESD22-B103B.01（振动，变频）或类似标准。
2.09 |  | Unit should be resistant to standard cleaning chemicals utilized in a hospital. | Unit should be cleanable by Quaternary Ammonium, Hypochlorite, Accelerated Hydrogen Peroxide, Phenolics, and Peracetic Acid without damage to casing or seals. | 可以用季铵盐、次氯酸盐、加速过氧化氢、酚醛和过氧乙酸清洗，不能损坏套管或密封件。
2.1 |  | Unit should be resistant to flame. | Casing, Gasket, PCB and other relevant components should carry non-flammable rating V-0. PCB | 外壳、垫片、PCB等相关部件应具有V-0不燃等级。
2.11 |  | Unit will need to operate in the wash cycle of an autoclave. | Unit will need to operate in a full wash cycle of an autoclave as specified below. The full wash is for 10.25 minutes. The Pre-wash for the first two minutes uses cold tap water. The Enzyme wash next is 1 minute at <40 C using a neutral enzymatic cleaner. The wash cycle is 5 minutes at 66 C using a neutral pH detergent. Then there are two rinse cycles at 2 minutes using > 40 C tap water and 0.25 minutes at approximately 22C using distilled water. | 设备将需要在如下所述的蒸压釜的完整洗涤循环中操作。 全洗时间为10.25分钟。 前两分钟的预洗用的是冷水。 接下来使用中性酶清洗剂在<40℃下洗1分钟。 使用中性pH洗涤剂，在66℃下洗涤5分钟。 然后在2min内用> 40℃自来水进行2次冲洗循环。在22℃用蒸馏水冲洗0.25分钟。
2.12 |  | Unit will need to operate in a thermal disinfection cycle of an autoclave. | Unit and electronics should be designed to withstand a thermal disinfection cycle of 6 minutes at >90 C using distilled water. | 装置和电子设备的设计应能承受使用蒸馏水在>90℃下6分钟的热消毒循环。
2.13 |  | Unit will need to operate in a drying cycle of an autoclave. | Unit and electronics should be designed to withstand an autoclave drying cycle of 7 minutes at 115 C air temperature using medical grade air. | 装置和电子设备的设计应能承受115℃空气温度下7分钟的高压灭菌器干燥循环，使用医用级空气。
2.14 |  | Unit will need to operate in a prevacuum cycle of an autoclave. | Unit and electronics should be designed to withstand a prevacuum cycle as specified below. The first phase has a minimum exposure time of 4-18 minutes at 132 C with a drying time of 30 minutes at temperatures less than 137C. The second phase has a minimum exposure time of 3-18 minutes at 134 C with a drying time of 30 minutes. | 设备和电子器件的设计应能承受以下规定的预真空循环。 第一阶段的最低暴露时间为4-18分钟，温度为小于137℃，干燥时间为30分钟。 第二阶段的最低暴露时间为3-18分钟，温度为134℃，干燥时间为30分钟。
3.01 | Electrical Requirements | Transmit power shall fit into Bluetooth Low Energy guidelines. | Transmit Power should be between 1 and 4 dBm | 发射功率应该在1到4 dBm之间 （蓝牙）
3.02 |  | Transmit distance shall meet industry norms. | Transmit Distance should be 70 m unobstructed | 发射距离应该是70米（无遮挡）
3.03 |  | Beacon shall be capable of measuring external temperature | Beacon will have a temperature sensing device capable of detecting from -25 C to 125 C.This temperature will be measured with every status packet. | （01.20细化需求）信标将有一个温度感应装置，能够检测到从-25℃到125℃的温度。这个温度将在每个状态数据包中测量。
3.03 |  | Beacon shall be capable of tracking battery percentage | Beacon will be able to measure the battery voltage and use this value to send a percentage.
3.04 |  | Beacon shall be able to track time between power cycle events. | Beacon will keep track of time since the last time it shut down from the temperature switch being triggered following the pattern listed in the packet specification.
3.05 |  | Beacon shall have enough non-volatile memory to store its firmware | Enough non-volatile memory will be included on the unit to provide enough space to store its firmware and a log of the last 50 power cycle events and the time that they occurred. | 包含足够的非易失性内存，以提供足够的空间来存储其固件和最近50个电源循环事件的日志以及它们发生的时间。
3.06 |  | Beacon shall be capable of turning itself off from an external stimulus and use this stimulus to activate its receiver for initial setup | Beacon will use a hall sensor that will trigger from a magnet touching the beacon for 10 seconds. This will activate the receiver for the next 12 hours to allow for updating the beacon firmware and its UUID, Minor and Major fields.
3.07 |  | Beacon shall shut off the receiver if the temperature switch is triggered. | Beacon will shut off the receiver if the temperature switch is triggered while it is active. The receiver will not be reenabled until the hall sensor is triggered again.
4.01 | Firmware Requirements | Beacon shall communicate using Bluetooth Low Energy to conserve power | Unit wil communicate using two iBeacon packets. One packet will contain the time since last sterilization and the other will contain the number of autoclave cycles that the unit has been through
4.02 |  | Unit shall record relevant informaiton for sterlization tracking. | Unit needs to broadcast four pieces of information: Location, number of lifetime cycles, battery percentage, and time since last sterilization.
4.03 |  | Unit should log power cycles to count autoclave cycles. | The unit should store up to 50 power cycle events with the uptime for each event. After the 50th event is recorded, a FIFO approach will be followed to discard the oldest data point and add the newest data point. | （01.20细化需求）该单元应存储多达 50 个电源循环事件以及每个事件的正常运行时间。 记录第 50 个事件后，将遵循 FIFO （先进先出）方法丢弃最老的数据点，并添加最新的数据点。
4.04 |  | Unit shall use iBeacon format for location information | The following fields should be included in the iBeacon1 Packet: 1) Data Length [1 byte - 0x02] 2) Data Type [1 byte - 0x01] 3) Value [1 byte - 0x06] 4) Length of the next section [1 byte - 0x1A] 5) Custom Manufacturer Data [1 byte - 0xFF] 6) Apple's Bluetooth Code [2 bytes - 0x004C] 7) Apple's iBeacon Indicator [2 bytes - 0x215] 8) Proximity UUID [16 bytes] 9) Major [2 bytes] 10) Minor [1 bytes] 11) Battery Indicator [1 bit] 12) Time Since Last Sterilization [7 bits] 13)Measured Power [2 bytes]
4.05 |  | Unit shall send two different iBeacon packets to allow for the transmission of all required information. | The following fields should be included in the iBeacon2 Packet: 1) Data Length [1 byte - 0x02] 2) Data Type [1 byte - 0x01] 3) Value [1 byte - 0x06] 4) Length of the next section [1 byte - 0x1A] 5) Custom Manufacturer Data [1 byte - 0xFF] 6) Apple's Bluetooth Code [2 bytes - 0x004C] 7) Apple's iBeacon Indicator [2 bytes - 0x215] 8) Proximity UUID [16 bytes] 9) Major [2 bytes] 10) Minor [1 bytes] 11) Second Battery Indicator [1 bit] 12) Total Number of Autoclave Cycles [7 bits] 13)Measured Power [2 bytes]
4.05 |  | Unit shall have defined fields for communication for its normal advertisement packets. | The following fields should be included in the Status Packet: 1) Data Length [1 byte - 0x1B] 2) Data Type [1 byte - 0x16] 3) UUID [2 bytes - 0x2BBB] 4) Sequence Number [4 bytes] 5) Device MAC Address [6 bytes] 6) Serial Number [6 bytes] 7) Internal Temperature [2 byte] 8) Battery Voltage [2 byte] 9) Current Uptime [4 bytes] | （02.11更新需求）以下字段应包含在状态包:1)数据长度(1个字节- 0 x1b) 2)数据类型(1个字节- 0 x16) 3) UUID(2字节- 0 x2bbb) 4)序列号(4个字节)5)设备的MAC地址(6字节)6)序列号(6字节)7)内部温度(2字节)8)电池电压(2字节)9)电流正常运行时间(4个字节)
4.06 |  | Unit shall have defined fields for communication with the status packet | The following fields shall be included in the History Packet: 1) Data Length [1 byte - 0x1B] 2) Data Type [1 byte - 0x16] 3)UUID [2 bytes - 0x2A14] 3)Sequence Number [2 bytes] 4) MAC Address [6 bytes] 5)Power Cycle Data [16 bytes - 4bytes per cycle] | （02.11更新需求）历史报文中包含以下字段:1)数据长度[1字节- 0x1B] 2)数据类型[1字节- 0x16] 3)UUID[2字节- 0x2A14] 3)序列号[2字节]4)MAC地址[6字节]5)重启[16字节-每周期4字节]
4.07 |  | Unit shall have a response packet enabled with defined fields | The following fields shall be included in the Response Packet: 1)Data Length [1 byte - 0x1D] 2)Data Type [1 byte - 0x16] 3)UUID [2 bytes -0x2A14] 4)Power Cycle Data [24 bytes - 4 bytes per cycle] | （02.11更新需求）响应报文中包含以下字段:1)数据长度[1字节- 0x1D] 2)数据类型[1字节- 0x16] 3)UUID[2字节- 0x2a14] 4)重启[24字节-每周期4字节]
4.08 |  | Beacon shall have the ability to send the full log of 50 power cycle events | Beacon will at minimum broadcase a History Packet containing 10 power cycles. A J&J application will be able to retrieve all 50 power cycle events. | （02.11更新需求）信标将至少广播一个包含10个功率循环的历史包。 强生的配套应用程序来检索最近50个功率循环的信息。
4.09 |  | Beacon should send a packet considering the status of the unit and a second larger packet with a log of the power cycles of the unit. | The Status packet should include all fields listed in 4.05. The History Packet should contain all fields listed in 4.06. The Response Packet should contain all fields listed in 4.07. | （01.20细化需求）Status packet应包括 4.05 中列出的所有字段。 History Packet应包含 4.06 中列出的所有字段Response Packet应包含 4.07 中描述的电源循环日志。
4.06 |  | Unit shall have a packet specification that details the information included in each packet as well as how that information is decoded. | Follow the latest packet specification (Currently 220407) for details on the different fields as well as how the information in the last sterilized, battery indicators, and total number of autoclave cycles is decoded and when those fields are reset.
4.07 |  | Beacon shall have the ability to perform an initial configuration to allow for updating firmware and initial setup of UUID, Major, and Minor fields. | Unit should have an initial period of time after the battery is first activated to allow for GATT communications to update the firmware if necessary and to set the UUID, Major, and Minor fields.
4.11 |  | Beacon Bluetooth advertisement interval shall be adjustable based on use case need. | A J&J application should be capable of setting the advertisement interval from 10 to 300 seconds for the Status Packet, set a separate interval from 10 to 300 seconds for the History Packet, and set a separate inteval from 10 to 300 seconds for the iBeacon Packet. | （01.20细化需求）强生的配套应用程序应该能够为Status Packet和History Packet，分别设置从 10 到 300 秒的广播间隔
4.08 |  | Beacon shall allow for firmware updates from a configuration application. | Beacon will have a handshaking procedure over bluetooth to allow for the firmware to be updated. | Beacon将有一个握手程序（通过蓝牙），以允许固件更新。
4.13 |  | UUID will be defined by the manufacturer | Standard packet will use the SIG UUID for Status flag [0x2BBB]. History packet and Response packet will use the SIG UUID for Reference Time Information [0x2A14]. | （02.11更新需求）状态数据包将使用SIG UUID作为状态标志[0x2BBB]。 历史包和响应包将使用SIG UUID作为参考时间信息[0x2A14]。
4.09 |  | Packets should be designed for Android and iOS compatibility | Both iBeacon packets will be configured to work with both Android and iOS systems.
4.1 |  | Ibeacon packets should be designed to be detectable by iOS when an app is in background mode | The iBeacon packet will need to be detectable if the application is in background mode.
4.16 |  | Sequence Number will increment when there is an important change | Status packets will only increment the sequence number when the internal temperature changes enough to go over the configured amount or the battery voltage changes. History packets will only increment the sequence number when there is a new power cycle event. | （01.20新增需求）仅当内部温度变化足以超过配置的量或电池电压发生变化时，状态数据包才会增加序列号。 历史数据包只会在有新的电源循环事件时增加序列号。
Note on D10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 电池我们确认设计是不可更换的。 The battery we confirm is designed to be non-replaceable Reply: Replaceable requirement has been removed.
Note on D12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 目前评估可做的理论温度范围上限是125℃（不完全保证，需要可靠性测试通过后确认） The upper limit of the theoretical temperature range currently evaluated as doable is 125°C (not fully guaranteed, needs to be confirmed after reliability tests are passed) Reply: Understood, this is acceptable.
Note on D21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please note the added clarity here to match the sterlization process document provided.
Note on D24: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Revised to 125C. Reply: Requirement removed.
Note on D33: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: MOKO: 1. iBeacon is a separate broadcast packet, which is completely separate from Status packet, History packet. So what is the default Advertising Interval of iBeacon and what is the range that can be set? 2. What are the contents of iBeacon, UUID, Major, Minor, whether they can be modified by GATT, and what are the requirements for the configuration of these parameters when mass production Reply: 1. Default Advertising Interval for iBeacon is 10 seconds with the idea that we are sending two different ones with five seconds between each. 2. Contents of all of the packets is included in the 220407 Packet Specirfication document. They should be initially modifiable by GATT and will be specified by the Packet Specification and further requests from JnJ.
Note on D34: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: MOKO: 1. iBeacon is a separate broadcast packet, which is completely separate from Status packet, History packet. So what is the default Advertising Interval of iBeacon and what is the range that can be set? 2. What are the contents of iBeacon, UUID, Major, Minor, whether they can be modified by GATT, and what are the requirements for the configuration of these parameters when mass production
Note on D35: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: These lines are removed from the requirements
Note on D38: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: MOKO提供可支持第三方APP（nRF Connect）的GATT通信协议用于功能验证。 适配产品的配置工具及APP由客户负责，MOKO不提供相关的开发和支持 MOKO provides a GATT communication protocol that supports third-party APPs (nRF Connect) for functional verification. The configuration tool and APP for the adapted product are the responsibility of the customer, MOKO does not provide development and support. Reply: We need MOKO to develop the companion APP or process to adjust the configurable fields listed in Requirement 4.09/4.10 (advertisement intervals) and 4.07 (retreive the full 50 power cycel device history).
Note on D39: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 这里是否写错？History Packet应包含 4.05中列出的所有字段以及 4.03 中描述的电源循环日志。 Is there a mistake here?? The History Packet should contain all the fields listed in 4.05 as well as the power cycle log described in 4.03. Reply: This was a typo, please see revisions.
Note on D44: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This is our current assumption, but needs to be verified based on iOS testing.
Note on D45: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Sample Packet是不是写错了？应该是Status packet Is Sample Packet written wrong? It should be Status packet Reply: Yes, this was a typo. Reply: Updated to remove the additional packet specifications
Note on D47: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Requirement 4.16 was a duplicate and has been deleted.